# page_15.png

UNCERTAINTY ANALYSIS OF GAS METRING SYSTEM

~

PM4 POWER PLANT

2.0 SYSTEM DESCRIPTION

The custody transfer gas metcring skid of Phu My GDC is Ultrasonic (USM) mctering

system.

Eacl systcmn consisís of two Ultrasonic flow metcrs, each sizcd for 100% Ilow 5 000 1n'/l

connected in serics with bypass lines around them on a skid. In normal use; one of them is used as

thne   operation

mcter   and the   other

as   standby

meler.  Normally   the

flow

shall

be  tlrough the

operation meler and shall bypass the slandby meler; by manually closing tle bypass line: During

opcralion meter maintcnancc; tlc standby meter shall be

as the

opcration mctcr.

Eacl

melcr

[UI

is   equipped

wilh

ohe

Danicl

FloBoss   $60o

flow   computer

[or

Ilow

computation

with

built

in   pressure

transmitter .

'Tempcralure

lransnitter

is

used

for

line

gas

(emperature measurement   with

a

(4-20) mA analoguc input into flow computer. The Daniel 500

Serics gas chromatograph (GC) is installed at thc Phu My GDC main control terminal for linc gas

composition   analyscs; GC'  data

is   distributed

into each

flow computer   via MODBUS  protocol

through RS-485 serial link. Flow cotnputcr shall  use the received gas data and normalizcd it lor

encrgy calculation.

Line   volume   (lowrale is caleulatcc in accordance with

AGA

9)

for ultrasonic  mnctcr. Line

compressibility is calculated in accordancc witl AGA 8. Standard volume flow is then derived with

the PIZ lormula and encrgy flow ratc is obtaincd with the calorilic value calculated  using ISO

6976.

Page 5 ( 31

used