# page_22.png

UNCERTAINTY ANALYSIS OF GAS METERING SYSTEM

"

PM4 POWER PLANT

4.0 STANDARD VOLUME FLOWRATE UNCERTAINTY

Uptol

4 0,050% of span for PT

1423

Uplol

=

+ 0,044% of span for PT = 1443

when

calculating the standard uncertainty a rectangle distribution is assumed with a coverage

factor k

=1,73 .

4.1.1.6 Drift Upd

The contribution due to drifì in transducer output is calculated from the value for

stability

taken from the

product data sheet and the calibration interval.

Stability

=4

0,125% URLI 60 months

As the upper range limit of tle transducer is 137,9 bar and the calibration interval

year, the

component assumed for drift is taken to be 0,125 % URL,

Upd = + 0,034 barl year

Once a suflicient number of calibrations have been performed over time this value can be

recalculaled from the calibration results.

when

calculating the standard uncertainty a rectangle distribution is assumed with a coverage

factor k = 1,73 .

4.1.1.7 Power supply effect

A conlribution is made lo the

uncertainty budget for variations in power supply, including the

variation   belween

the power  supply   during  the   calibration

of  the

transmitter

and   that   during

operation:. he estimated uncerlainty for this is taken from the product data sheet.

Power supply effect

=

<

4

0,005% of calibrated span per volt

The tolerance for the maximum power supply deviation is assumed t0 be + 5 volts, therefore

=

4 0,025% of calibrated span

When   calculating   the

standard   uncertainty

à

rectangular   distribution

is

assumed

with

a

coverage factor k = 1,73 .

4.1.1.8

Combined uncertainty

The combined standard uncertainty calculated using the equation

Ptc

pcal

ptemp

pude

Um

4

V3

3

is represented in a tabular form as shown below; note that the sensitivity coellicient ci for each input

is 1

as

are all expressed as a percentage of the pressure reading.

Page 12 ol 31

Upp

Up"7

ptol

M p

they