# page_23.png

UNCERTAINTY ANALYSIS OF GAS METERING SYSTEM

"

PM4 POWER PLANT

4.0 STANDARD VOLUME FLOWRATE UNCERTAINTY

PT

~

1423

Source

Value

Symbol

Distribulion

[ivisor

C;

Ilj

(ci x u;)?

of Uncertainty

"

Upac

Reference accuracy

0,026

Rectangular

1,73

1,00

0,015

2,341-04

Upcal

Transmitter calibration

0,061

Normal

2,00

1,00

0,030

9,1BE-04

Ulemp

Ambient_tenperature_eflcct

0,059

Rectangular

1,73

1,00

0,034

1,47E-03

Uadc

ADC tolerance (llowConp_Input)

0,09[

Kectangular

1,73

1,00

0,052

2,75E-03

Uptol

'[ransmitter calibration tolerance

0,091

Keclangular

1,73

1,00

0,052

2,75E-03

Upd

Drift

0,104

Reclangular

1,73

1,00

0,060

3,64E-03

Power

elfect

0,045

Rectangular

1,73

1,00

0,026

6,89E-04

u(p)

Combined standard ucertainty

Nornnal

0,110

U(p

Fxpanded uncertainty

Normal (k-2)

2,00

05322

PT _ 1443

Source

Value

Synhol

Distrilulion

[ivisor

Ci

U;

(c; X u;)?

of Uncertainty

%

Upac

Reference accuracy

0,025

Rectangular

1,73

1,,00

0,014

2,0972 04

Upcal

[ransmitter calibration

0,/14

Normal

200

L0o

9,057

3,20

Utemp

Ambient_temperature elfect

0,043

Rectangular

173

1,0o

0,025

6,21E-04

Uade

ADDC tolerance (ElowConp_Input)

0,091

Kectangular

1,73

1,00

0052

2,75E-03

Uptol

Transmitter calibralion tolerance

0,080

Reclangular

1,73

1,00

0,046

2,11E-03

Upd

Urift

0,209

Rectangular

1,73

1,00

0,121

1,46E-02

Upp

Power

supply effect

0,045

Rectangular

1,73

1,00

0,026

6,89E-04

"(p)

Combined standard uncertainty

Normal

0,156

U(j)

Expandled uncertainty

Normal (k=2)

2,00

0,31;

4.1.1.9 The Expandcd uncertainty

The expanded unccrtainty is calculatcd by multiplying the combined standard  unccrtainty hy

thc coverage lactor k

2

2.

4.1.2 Temperature

'Temperalure is mcasurcd

a RID PTIO0.

'Thc  output is a resistancc  analoguc signal

ranged (0

80) %'

The combined standard  uncertainty in tcmpcraturc is cslimated by combining inputs dcrivcd

{rou the  calibralion of thc  transduccr   and

thc  mnanufacturer  $

fication

thc

following

equation:

Pagc 13 0f 31

supply

Upe

using

spcci

using