# page_26.png

UNCERTAINTY ANALYSIS OF GAS METERING SYSTEM

~

PM4 POWER PLANT

4.0 STANDARD VOLUME FLOWRATE UNCERTAINTY

Utp

4 0,18% of reading

when calculating the slandard uncertainty a rectangle distribution is assumed with a coverage

factor k

=

1,73 .

4.1.2.8 ADC tolerance

padlc

The

analogue to digital converter at the input of the {low computer is checked by routine

calibration to be within the tolerance of + 0,1% of span. However the

after calibration is:

=+

0,01 % of reading for TT

1424

=+

0,01 % of reading for TT

1444

When

calculating

the

standard   uncertainty

a

rectangular   distribution

is

assumed

with

a

coverage factor k = 1,73 .

4.1.2.9 Combined uncertainty

The combined standard uncertainty calculaled

the equation

Adac:

teal

llemp

Itul

Jaudlc

U,

2

is represented in a tabular form as shown below, note that the sensitivity coefficient c; lor each

input is 1

as

are all expressed as a percentage ot the lemperature reading.

For TT _ 1424

Source

Value

Symbol

Distribution

Divisor

Ci

Ili

(c; X v;)?

of Uncertainty

Digital accuracy_

0,500

Rectangular

1,73

1,00

0,289

8,35E-02

DIA_Accuracy

0,080

Rectangular

1,73

1,00

0,046

2,14E-03

Transmitter calibration

0,700

Normal

2,00

1,00

0,350

1,23E-01

Uuenp

Ambient temperature effect

0,050

Rectangular

1,73

1,00

0,029

8,336-04

Transmitter calibration

Uuol

tolerance

0,/90

Rectangular

1,73

1,00

02110

1,20E-02

Uu

Drift

0,050

Rectangular

1,73

1,00

0,029

8,33E-04

Power

effect

0,100

Rectangular

1.73

1,00

0,058

3,33E-03

Uladc

ADC test

0,040

Rectangular

1,73

1,00

0,023

5,331-04

U()

Combined standard_uncertainty

Normal

0,475

Uo

Fxpanded uncertainty

Normal (k=2)

2,00

0,95

Page 16 of 31

Uudac <

Uadc

Utadc

using

+(%;(

they

Uudlac

Uudu

Ucal

Uw

supply