# page_27.png

UNCERTAINTY ANALYSIS (F GAS METERING SYSTEM

PM4 POWER PLANT

4.0 STANDARD VOLUME FLOWRATE UNCERTAINTY

For TT

~

1444

Source

Value

Symhol

Distribution

Divisor

Ci

U;

(ci X u;)?

of Uncertainly

%

Digital accuracy

0,500

Rectangular

1,73

1,00

0,289

8,35E-02

D/A_Accuracy

0,080

Rectangulax

73

1,00

0,046

2,14E-03

Transmitter calibration

0,700

Normal

2,00

1,00

0,350

1.23E-01

Uewj

Ambient (emperaturc c(lect

0,050

Rcctangular

1,73

1,00

0,029

8,33E-04

Transmitler calibration

lolerance

0,250

Rectangular

1,73

1,00

0,144

2,081-02

Uud

Drilt

0,050

Rectangular

1.73

1,00

0,029

8,33E-04

Ulp

Power_supply_effect

0,100

Rectangular

1,73

1,00

0,058

3,33E-03

Ulade

ADC' lest

0,040

Rectangular

1,73

1,00

0,023

5,33E-04

Combined standard_unccrtainty

Normal

0.484

U

Expanded uncertainty

Normal (k=2)

2.00

0,97

4.1.2.16) Thc Expanded uncertainty

The expandcd unccrtainty is calculated by multiplying tluc combincd standard unccrlainty by

thc coverage lactor k

=

2.

4.1.3 (as composition

The   composition ol thne  gas is mcasurcd by

a

gas   chromatograph

whicl is calibraled by

exposure (0 a certificd calibration gas witl a

composition rcpresentalivc of the gas to be melered.

Ile calibration gas comprises the lollowing componcnts:

Nilrogen

Nz

Carbon Dioxide

Mcthanc

C1

Fthane

C2

Propane

C3

nl3ulane

IC4

iBulanc

iC4

nentane

nCs

iPentanc

iCs

neoPcntane

ncoCs

Ilcxane

nC

The combined standard unccrtainty [or cach component is calculaled by

U

calg

tol

rep

U mn :=

17 of 31

Uudac

Utda

Ucal

Uugl _

U(u)

C0z

Page