## Page: page_15.png

compressibility is calculated in accordancc witl AGA 8. Standard volume flow is then derived with

the PIZ lormula and encrgy flow ratc is obtaincd with the calorilic value calculated using ISO

UNCERTAINTY ANALYSIS OF GAS METRING SYSTEM ~ PM4 POWER PLANT

2.0 SYSTEM DESCRIPTION

The custody transfer gas metcring skid of Phu My GDC is Ultrasonic (USM) mctering

system.

Eacl systcmn consisís of two Ultrasonic flow metcrs, each sizcd for 100% Ilow 5 000 1n'/l

connected in serics with bypass lines around them on a skid. In normal use; one of them is used as

thne   operation mcter   and the   other as   standby meler.  Normally   the flow shall be  tlrough the

operation meler and shall bypass the slandby meler; by manually closing tle bypass line: During

opcralion meter maintcnancc; tlc standby meter shall be used as the opcration mctcr.

Eacl melcr [UI is   equipped wilh ohe Danicl FloBoss   $60o flow   computer [or Ilow

computation with built in   pressure transmitter . 'Tempcralure lransnitter is used for line gas

(emperature measurement   with a (4-20) mA analoguc input into flow computer. The Daniel 500

Serics gas chromatograph (GC) is installed at thc Phu My GDC main control terminal for linc gas

composition   analyscs; GC'  data is   distributed into each flow computer   via MODBUS  protocol

through RS-485 serial link. Flow cotnputcr shall  use the received gas data and normalizcd it lor

encrgy calculation.

Line   volume   (lowrale is caleulatcc in accordance with AGA 9) for ultrasonic  mnctcr. Line

6976.

Page 5 ( 31