## Page: page_17.png

| p | Gas density al linc condition; calculated in accordancc will AGA 8 - 1994 |
| --- | --- |
| 0l | Gas density at base condition, calculated in accordancc with IS0 6976-1995 |

| 5,0 | 0,15 | o,000 | 0,30 | 0,14 | 0,29 |
| --- | --- | --- | --- | --- | --- |
| 10,! | 017 | u,uuo | 0,30 | 0,15 | 0,30 |
| 25,2 | 0,217 | @,u0o | 0,30 | 0215 | 0,30 |
| 39,9 | 0,15 | @,000 | ",30 | 014 | 0,29 |
| 69,8 | 0,14 | (),uo0 | 0,30 | 044 | (,28 |
| 80,4 | 015 | Q,uou | 0,30 | 0,14 | 0,29 |

UNCERTAINTY ANALYSIS OF GAS METERING SYSTEM PM4 POWER PLANT

3.0 ULTRASONIC METER UNCERTAINTY

1

Pressure al linc condition. bar a

1

Prcssurc at basc condition; 1,01325 bar a

T,

Tcmperalurc at line condition; %C'

T;

'Temperalure at basc condition; 15 %C (288,15 K)

Z,

Gas compressibility at line condition; calculated in accordancc wilh AGA 8-1994

Z,

Gas compressibility at basc condition; calculalcd in accordance witl IS0 6976-1995

3.3 Derivation of ultrasonic metering uncertainty

The  ultrasonic  tnelcr was senl lor 2 wet  calibration at a third party calibration facility in

C'EESI. From thc calibralion certificale; CILSI had dcclared that Ile mnaximum tolerance crror of

the calibrated mcter is (,064 % for FT 1420. 0,05 % lor FT 1440.

From the experimcntal rcsult ol the calibration; the maximum calibration uncertainty (Ucal) of

melers which will allcct the actual mncasurcmnent during opcralion is estimated for this experimental

dala 0,18 % ol calibrated span.

An installation conliguration error will contributed to the meter installalion at site (Umi).

9 Section 7.2.2. recommended that a maximun error is 0,3 % lor this effect and it has accounted for

the k multiplicr .

Thne USM combined uncertainty is givcn a lollow

cul Uol mi

U(m) =

2 1,73 2,45

Thic calculated cxpandcd unccrtainty lor ultrasonic meter is summarized in tables below

For FT _ 1420

Standard uncertainty (% " Combined Expanded

Ultrasonic meter

uncertainty

flow rangc (%) uncertainty

Ucal Utol Umi k = 2

ci 1,(0 1,00 1,00

Pngc 7 0f31