## Page: page_21.png

UNCERTAINTY ANALYSIS OF GAS METRING SYSTEM PM4 P(WIR PLANT

4.0 STANDARD VOLUME FLOWRATE UNCERTAINTY

when calculaling Ilc slandard uncertainty a rcctangle distribution is assumed witl a coverage

factor k = 1,73.

4.1.1.2 Transducer calibration U pcal

The pressure transduccrs arc calibrated at site, witl all corrcctions madc it is estimated that

the unccrtainly in measurement is calculatcd lron calibration reports at site and associated witl the

rclerence slandard uncertainty . So

Upcal 1 0,033% of span for PT 1423

Upcal 4 0,063% of span lor PT = 1443

when calculating the slandard uncertainty a normal distribution is assumed with 2 coveragc

faclor k = 2

4,1.1.3 Amhient temperaturc elfect Uptemp

conlribution is nade l0 thc  uncerlainty   budget for variations in ambient   temperalure,

including the variation betwcen the tcmperalure  during the calibration of thc transducer and tlic

opcraling tempcraturc . The estimated uncertainty for this is taken [ron the product data sheet.

For turndown ratio for span grcatcr than I/10 of URL

Uplenp = 4 (0,0125% URL + 0,0625% span) over 28  C

As the URL of transducer is 137,9 bar and the span is 60 bar, the resulting in an uncer

of:

Uptcmp = 4 0.055 bar/28 %C

The tolerance for the maximun temperature deviation is assumed 10 be + 10  C, therelore the

uncertainty duc l0 tcmperature changes is

Upemp = + (.049 % of reading

when   calculating thc slandard   uncertainty à rcclangle distribution is   assumed with a

coverage factor k = 1.73 .

4.1.1.4 Al)C' tolerance padlc

The  analoguc to digital converler at the input of the  flow computer is chccked by rouline

calibration t0 be within the tolerance o[+ 0,1% of span. However the Updac altcr calibration is:

Upade = 4 0,05 % of reading for PT - 1423

Upadlc =2+ 0,04 % of rcading lox PT 1443

Whcn   calculaling lhe slandard   unccrtainty à rcclangular   distribution is assumcd witli à

coverage factor k = 1,73.

4.1.1.5 Transduccr calilration tolcrancc Uptol

[he acceptance crilerion lor the calibration of thc prcssurc (ransducers is equal to the absolule

maximnumn error of test points ovcr calibrated span. Therefore

Tage 11 0f 31