## Page: page_22.png

component assumed for drift is taken to be 0,125 % URL,

Upd = + 0,034 barl year

UNCERTAINTY ANALYSIS OF GAS METERING SYSTEM " PM4 POWER PLANT

4.0 STANDARD VOLUME FLOWRATE UNCERTAINTY

Uptol 4 0,050% of span for PT 1423

Uplol = + 0,044% of span for PT = 1443

when calculating the standard uncertainty a rectangle distribution is assumed with a coverage

factor k =1,73 .

4.1.1.6 Drift Upd

The contribution due to drifì in transducer output is calculated from the value for stability

taken from the product data sheet and the calibration interval.

Stability =4 0,125% URLI 60 months

As the upper range limit of tle transducer is 137,9 bar and the calibration interval year, the

Once a suflicient number of calibrations have been performed over time this value can be

recalculaled from the calibration results.

when calculating the standard uncertainty a rectangle distribution is assumed with a coverage

factor k = 1,73 .

4.1.1.7 Power supply effect Upp

A conlribution is made lo the uncertainty budget for variations in power supply, including the

variation   belween the power  supply   during  the   calibration of  the transmitter and   that   during

operation:. he estimated uncerlainty for this is taken from the product data sheet.

Power supply effect = < 4 0,005% of calibrated span per volt

The tolerance for the maximum power supply deviation is assumed t0 be + 5 volts, therefore

Up"7 = 4 0,025% of calibrated span

When   calculating   the standard   uncertainty à rectangular   distribution is assumed with a

coverage factor k = 1,73 .

4.1.1.8 Combined uncertainty

The combined standard uncertainty calculated using the equation

Ptc pcal ptemp pude ptol Um

M p 4

V3 3

is represented in a tabular form as shown below; note that the sensitivity coellicient ci for each input

is 1 as they are all expressed as a percentage of the pressure reading.

Page 12 ol 31