## Page: page_23.png

| Uade | ADDC tolerance (ElowConp_Input) | 0,091 | Kectangular | 1,73 | 1,00 | 0052 | 2,75E-03 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Uptol | Transmitter calibralion tolerance | 0,080 | Reclangular | 1,73 | 1,00 | 0,046 | 2,11E-03 |

UNCERTAINTY ANALYSIS OF GAS METERING SYSTEM " PM4 POWER PLANT

4.0 STANDARD VOLUME FLOWRATE UNCERTAINTY

PT ~ 1423

Source Value

Symbol Distribulion [ivisor C; Ilj (ci x u;)?

of Uncertainty "

Upac Reference accuracy 0,026 Rectangular 1,73 1,00 0,015 2,341-04

Upcal Transmitter calibration 0,061 Normal 2,00 1,00 0,030 9,1BE-04

Ulemp Ambient_tenperature_eflcct 0,059 Rectangular 1,73 1,00 0,034 1,47E-03

Uadc ADC tolerance (llowConp_Input) 0,09[ Kectangular 1,73 1,00 0,052 2,75E-03

Uptol '[ransmitter calibration tolerance 0,091 Keclangular 1,73 1,00 0,052 2,75E-03

Upd Drift 0,104 Reclangular 1,73 1,00 0,060 3,64E-03

Upe Power supply elfect 0,045 Rectangular 1,73 1,00 0,026 6,89E-04

u(p) Combined standard ucertainty Nornnal 0,110

U(p Fxpanded uncertainty Normal (k-2) 2,00 05322

PT _ 1443

Source Value

Synhol Distrilulion [ivisor Ci U; (c; X u;)?

of Uncertainty %

Upac Reference accuracy 0,025 Rectangular 1,73 1,,00 0,014 2,0972 04

Upcal [ransmitter calibration 0,/14 Normal 200 L0o 9,057 3,20

Utemp Ambient_temperature elfect 0,043 Rectangular 173 1,0o 0,025 6,21E-04

Upd Urift 0,209 Rectangular 1,73 1,00 0,121 1,46E-02

Upp Power supply effect 0,045 Rectangular 1,73 1,00 0,026 6,89E-04

"(p) Combined standard uncertainty Normal 0,156

U(j) Expandled uncertainty Normal (k=2) 2,00 0,31;

4.1.1.9 The Expandcd uncertainty

The expanded unccrtainty is calculatcd by multiplying the combined standard  unccrtainty hy

thc coverage lactor k 2 2.

4.1.2 Temperature

'Temperalure is mcasurcd using a RID PTIO0. 'Thc  output is a resistancc  analoguc signal

ranged (0 80) %'

The combined standard  uncertainty in tcmpcraturc is cslimated by combining inputs dcrivcd

{rou the  calibralion of thc  transduccr   and thc  mnanufacturer  $ spcci fication using thc following

equation:

Pagc 13 0f 31