## Page: page_24.png

| Uulac | Digital accuracy |
| --- | --- |
| Uuda | DIA Accuracy |

UNCERTAINTY ANALYSIS OF GAS METERING SYSTEM " PMA POWER PLANT

4.0 STANDARD VOLUME FLOWRATE UNCERTAINTY

Iduc Ical Item Ilol tade'

U,

V3 3 v3 V3

where the sources of uncertainty are

Ical Transducer calibration

'Itetlp Ambient temperature effect

Ilu] Transducer calibration tolerance

Uu Drift

Uip Power supply effect

Uuadc ADC tolerance

4.1.2.1 Digital accuracy Uuac

The reference digital accuracy is taken from the product data sheet for temperature RTD of

flowcomputer .

Uulae =10,1 %C

when calculating the standard uncertainty a rectangle distribution is assumed with a coverage

factor k = 1,73.

4.1.2.2 D/A accuracy Utda

The reference D/A accuracy is taken lrom the product data sheed of RTD

Uda= = 4 0,02% of the calibrated span

The data sheet presents the specifications "conforming to at least 3 ơ" therefore when calculaling

the standard uncertainty a rectangle distribution is assumed witl a coverage lactor k = 1,73.

4.1.2.3 Transducer calibration Uucal

The temperature transmitler are calibrated at lab. with all corrections made it is estimated thal

the uncertainty in measurement associated witl the reference standard S0

cal 4 0,175 % of span lor TI' 1424

tcal I 0,175 % of span for TT 1444

when calculating the standard uncertainty a normal distribution is assumed with a coverage

3 Rosemount Model 3144P Temperature Transmitter

Page 14 of 31