## Page: page_27.png

| Uudac | Digital accuracy | 0,500 | Rectangular | 1,73 | 1,00 | 0,289 | 8,35E-02 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Utda | D/A_Accuracy | 0,080 | Rectangulax | 73 | 1,00 | 0,046 | 2,14E-03 |

| Uugl _ | lolerance | 0,250 | Rectangular | 1,73 | 1,00 | 0,144 | 2,081-02 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Uud | Drilt | 0,050 | Rectangular | 1.73 | 1,00 | 0,029 | 8,33E-04 |

| Mcthanc | C1 |
| --- | --- |
| Fthane | C2 |
| Propane | C3 |
| nl3ulane | IC4 |
| iBulanc | iC4 |
| nentane | nCs |
| iPentanc | iCs |

UNCERTAINTY ANALYSIS (F GAS METERING SYSTEM PM4 POWER PLANT

4.0 STANDARD VOLUME FLOWRATE UNCERTAINTY

For TT ~ 1444

Source Value

Symhol Distribution Divisor Ci U; (ci X u;)?

of Uncertainly %

Ucal Transmitter calibration 0,700 Normal 2,00 1,00 0,350 1.23E-01

Uewj Ambient (emperaturc c(lect 0,050 Rcctangular 1,73 1,00 0,029 8,33E-04

Transmitler calibration

Ulp Power_supply_effect 0,100 Rectangular 1,73 1,00 0,058 3,33E-03

Ulade ADC' lest 0,040 Rectangular 1,73 1,00 0,023 5,33E-04

U(u) Combined standard_unccrtainty Normal 0.484

U Expanded uncertainty Normal (k=2) 2.00 0,97

4.1.2.16) Thc Expanded uncertainty

The expandcd unccrtainty is calculated by multiplying tluc combincd standard unccrlainty by

thc coverage lactor k = 2.

4.1.3 (as composition

The   composition ol thne  gas is mcasurcd by a gas   chromatograph whicl is calibraled by

exposure (0 a certificd calibration gas witl a composition rcpresentalivc of the gas to be melered.

Ile calibration gas comprises the lollowing componcnts:

Nilrogen Nz

Carbon Dioxide C0z

neoPcntane ncoCs

Ilcxane nC

The combined standard unccrtainty [or cach component is calculaled by

U tol

calg rep

U mn :=

Page 17 of 31